| Code | Description |  | Examples |
| --- | --- | --- | --- |
| STA | Read Device Status – Reads the internal status of the laser. The response is the command echoed back, a delimiter of “: “, and then the status as a bit-encoded 32-bit unsigned word. Undefined bits or bits defined as “Reserved” can be in any state and should be ignored. Each of the bits has the following meaning: |  | Sent: “STA” Response: “STA: 4100” This translates to the following: 4100 = 0x1004, so bits 2 and 12 are set. This means that emission is on and modulation is enabled. |
|  | Bit 0: | Reserved. |  |
|  | Bit 1: | 0 = Temperature OK. 1 = Overtemperature Condition. |  |
|  | Bit 2: | 0 = Emission is Off. 1 = Emission is on or starting. |  |
|  | Bit 3: | 0 = Backreflection Level is OK. 1 = High Backreflection Level. |  |
|  | Bit 4: | 0 = Analog Control Mode is disabled. 1 = Analog control mode is enabled. |  |
|  | Bit 5: | Reserved. |  |
|  | Bit 6: | 0 = All modules communicating. 1 = Module(s) disconnected. |  |
|  | Bit 7: | 0 = All modules OK. 1 = Module(s) failed. |  |
|  | Bit 8: | 0 = Aiming Beam Off. 1 = Aiming Beam On. |  |
|  | Bit 9: | Reserved. |  |
|  | Bit 10: | Reserved. |  |
|  | Bit 11: | 0 = Power Supply OK. 1 = Power Supply Off. |  |
|  | Bit 12: | 0 = Modulation Disabled. 1 = Modulation Enabled. |  |
|  | Bit 13: | Reserved. |  |
|  | Bit 14: | Reserved. |  |
|  | Bit 15: | 0 = Emission is not in the 3 second safety start-up state. 1 = Emission is in the 3 second safety start-up state. |  |
|  | Bit 16: | Reserved. |  |
|  | Bit 17: | 0 = Control OK 1 = Unexpected emission detected. |  |
|  | Bit 18 – Bit 31: | Reserved. |  |