## Slide 1
areaDetector: A module for EPICS area detector support
Mark Rivers
GSECARS, Advanced Photon Source
University of Chicago

## Slide 2: areaDetector Talk Outline
Introductory talk
Motivation & goals for areaDetector module
Overview of architecture
Drivers for detectors & cameras
Plugins for real-time processing
Viewers and other clients
Demo with simDetector and FLIR USB-3 camera

## Slide 3: areaDetector - Goals
Drivers for many detectors, especially those used at synchrotron beamlines
  Handle detectors ranging from >500 frames/second to <1 frame/second
Basic parameters for all detectors
  E.g. exposure time, start acquisition, etc.
  Allows generic clients to be used for many applications
Easy to implement new detector
  Single device-driver C++ file to write. EPICS independent.
Easy to implement detector-specific features
  Driver understands additional parameters beyond those in the basic set
Middle-level plug-ins to add capability like regions-of-interest calculation, file saving, etc.
  Device independent, work with all drivers
  Below the EPICS database layer for highest performance

## Slide 4
NDArray
  N-Dimensional array.
    Everything is done in N-dimensions (up to 10), rather than 2. This is needed even for 2-D detectors to support color.
    Other types of devices (Xspress3 and xMAP x-ray spectrometers, quad electrometers also use NDArrays and areaDetector plugins.
  This is what plug-ins callbacks receive from device drivers.
NDAttribute
  Each NDArray has a list of associated attributes (metadata) that travel with the array through the processing pileline. Attributes can come from driver parameters, any EPICS PV, or any user-written function.
    Can store motor positions, temperature, ring current, amplifier gains, etc. with each frame.
    Written to disk files for TIFF, netCDF, and HDF5 file formats.
NDArrayPool
  Allocates NDArray objects from a freelist
  Plugins access in readonly mode, increment reference count
  Eliminates need to copy data when sending it to callbacks.
areaDetector – Data structures

## Slide 5
ADBase .template
Standard asyn device support (device-independent)
Vendor API
Driver
Channel Access Clients (CSS, medm, Python, ImageJ, SPEC, etc.)
EPICS areaDetector Architecture
xxxDriver .template
Hardware
C++ Base classes (NDArray, asynPortDriver, asynNDArrayDriver, ADDriver, NDPluginDriver)
Layer 5 Standard EPICS records
Layer 4
EPICS device support
Layer 3 Plug-ins
Layer 1
Hardware API
Layer 2
Device drivers
Layer 6
EPICS CA clients
StdArrays
File (HDF5, netCDF, TIFF, JPEG)
NDPluginXXX. template
Channel access
Record/device support
asynInt32, Float64, Octet
asynGenericPointer (NDArray)
asynXXXArray
C library calls
NDPluginBase .template
ROI
Process

## Slide 6: Detector Drivers
Currently ~35 detector drivers covering a wide variety of detectors.
Simulation driver
GigE cameras (Prosilica/AVT, Point Grey/FLIR, any GigEVision camera via aravis library)
FLIR and AVT USB-3.x cameras
Dectris Pilatus and Eiger pixel array detectors
Princeton Instruments and Photometrics detectors and spectrometers
Andor CCD and CMOS cameras
Perkin Elmer and Dexela flat panel detectors
Web cameras and Axis video servers
Many more (Basler, Pixirad, Photonic Sciences, etc.)

## Slide 7: ADBase.adl – Generic control screen
Works with any detector
Normally write custom control for each detector type to hide unimplemented features and expose driver-specific features

## Slide 8: simDetector: Detector-specific screen
Example
1024x1024 pixels
16-bit integer images
485 frames/s
~1GB/s

## Slide 9: simDetector Driver
3 simulation modes, LinearRamp, Peaks, Sine Waves.
Independent gains for X, Y, Overall, Red, Green, Blue
Linear ramp has X and Y linear ramp with array index, each cycle just adds value to each pixel. Very fast.
Peak mode is array of Gaussian peaks plus noise. Slower.
Sine mode is 2 sine waves in each of Y and Y, summed or multiplied. Slower.

## Slide 10: simDetector: Linear Ramp Mode

## Slide 11: simDetector Peaks mode with FFT

## Slide 12: simDetector: Sine mode

## Slide 13: Pilatus specific control screen

## Slide 14: LightField driver

## Slide 15: LightField driver

## Slide 16: URL Driver
Driver that can read images from any URL.
Can be used with Web cameras and Axis video servers.
Uses GraphicsMagick to read the images, and can thus handle a large number of image formats (JPEG, TIFF, PNG, etc.).

## Slide 17: Perkin Elmer Flat Panel Driver

## Slide 18: FLIR/Point Grey drivers
ADSpinnaker driver for all cameras from FLIR/Point Grey using their Spinnaker SDK.
GigE, USB 3.0, and 10 GigE camera
High performance, low cost
I will demonstrate USB-3.0 camera today
Example models:
Model | Interface | Resolution | Price | Measured Speed (frames/s) | Measured Speed (MB/s)
BlackFlyS BF2-PGE-13Y3M | 1-Gbit Ethernet | 1280x1024 | $325 | 83 frames/s | 104 MB/s
Grasshopper 3 GS3-U3-23S6M | USB-3 | 1920x1200 | $995 | 156 frames/s | 343 MB/s
Oryx ORX-10G-51S5M | 10-Bbit Ethernet | 2448x2048 | $2,258 | 163 frames/s | 779 MB/s

## Slide 19: Point Grey GigE Camera BlackFly PGE-20E4C
e2v EV76C570 CMOS sensor
Global shutter
29 x 29 x 30 mm
Power Over Ethernet
4.5 micron pixels
1600 x 1200 pixels, color or mono)
50 frames/s
$525

## Slide 20: Point Grey USB-3.0 Camera Grasshopper3 GS3-U3-23S6M
1920 x 1200 global shutter CMOS
Sony IMX174 1/1.2
Dynamic range of 73 dB
Peak QE of 76%
Read noise of 7e-
12-bit or 8-bit data
Max frame rate of 162 fps
  ~356 MB/S, >3X faster than GigE
USB 3.0 interface
Used for tomography at 3 APS beamlines, replaced Andor Neo and PCO Edge
$995

## Slide 21: Point Grey 10-Gbit Ethernet Camera Oryx ORX-10G-310S9M
6464 x 4852 global shutter CMOS (31 MP)
Sony IMX342 3.45 micron pixels
Dynamic range of 66 dB
Peak QE of 65%
Read noise of 5.2e-
12-bit, 10-bit, or 8-bit data
Max frame rate of 26 fps
  778 MB/S, >8X faster than GigE
$7,089
Used for tomography on 2 APS beamlines
Model | Resolution | Price | Speed (frames/s) | Speed (MB/s)
ORX-10G-310S9M | 6464x4852 | $7,089 | 26 frames/s | 778 MB/s
ORX-10G-123S6M | 4096x3000 | $4,666 | 68 frames/s | 797 MB/s
ORX-10G-51S5M | 2448x2048 | $2,258 | 162 frames/s | 774 MB/s
ORX-10GS-34S4M | 2048x1536 | $1,774 | 216 frames/s | 648 MB/s

## Slide 22: ADSpinnaker Driver

## Slide 23: Camera-specific Features Screen (2/6)
Database file and OPI screens built from GenICam camera XML downloaded from the camera using Python script

## Slide 24: Plugins
Designed to perform real-time processing of data, running in the EPICS IOC (not over EPICS Channel Access)
Receive NDArray data over callbacks from drivers or other plugins
Plug-ins can execute in their own threads (non-blocking) or in callback thread (blocking)
  If non-blocking then NDArray data is queued
  If executing in callback thread, no queuing, but slows driver
Allows
  Enabling/disabling
  Throttling rate (no more than 0.5 seconds, etc)
  Changing data source for NDArray callbacks to another driver or plugin
Plugins can be sources of NDArray callbacks, as well as consumers
  Allows creating a data processing pipeline running at very high speed, each in a different thread, and hence in multiple cores on modern CPUs.

## Slide 25: NDPluginDriver medm Screens

## Slide 26: Plugins (continued)
Currently 20 plugins that perform wide variety of operations
NDPlugInStdArrays
  Receives arrays (images) from device drivers, converts to standard arrays, e.g. waveform records.
  This plugin is what EPICS channel access viewers normally talk to.
NDPluginPVA
  Converts NDArrays to EPICS V4 NTNDArrays
  Exports the NtNDArrays over PVAccess with internal V4 server
  Can be used to send structured data to EPICS V4 clients
  When used with the PVAccess driver then areaDetector plugins can be run on different machine from the detector driver
NDPluginROI
  Performs region-of-interest calculations
  Select a subregion. Optionally bin, reverse in either direction, convert data type.
  Divide the array by a scale factor, which is useful for avoiding overflow when binning.
NDPluginTransform
  Performs geometric operations (rotate, mirror in X or Y, etc.)

## Slide 27: Plugins (continued)
NDPluginStats
  Calculates basic statistics on an array (min, max, sigma)
  Optionally computes centroid centroid position, width and tilt.
  Optionally Computes X and Y profiles, including average profiles, profiles at the centroid position, and profiles at a user-defined cursor position.
  Optionally computes the image histogram and entropy
NDPluginROIStat
  Multiple ROIs with simple statistics in a single plugin
  More efficient when many ROIs are needed, e.g. for peaks in a 1-D energy spectrum
  Min, max, total, net, mean
  Time-series of each of these statistics

## Slide 28: Plugins (continued)
NDPluginProcess
  Does arithmetic processing on arrays
  Background subtraction.
  Flat field normalization.
  Offset and scale.
  Low and high clipping.
  Recursive filtering in the time domain.
  Conversion to a different output data type.
NDPluginOverlay
  Adds graphic overlays to an image.
  Can be used to display ROIs, multiple cursors, user-defined boxes, text, etc.
ffmpegServer
  MJPEG server that allows viewing images in a Web browser. From DLS.

## Slide 29: Plugins (continued)
NDPluginAttribute
  Extracts NDAttributes from NDArrays and publishes their values as ai records
  Can collect time-series arrays of the attribute values
NDPluginCircularBuff
  Buffers NDArrays in a circular buffer
  Computes a trigger expression using up to 2 NDAttribute values
  When trigger condition is met then outputs NDArrays
  User-specified number of pre-trigger and post-trigger arrays to output
NDPluginTimeSeries
  Accepts 1-D NDArrays[NumSignals] or 2-D [NumSignals,NewTimePoints] and appends to time-series buffer
  Operates in fixed length (stop when full) or circular buffer modes
  Optional time-averaging of input data

## Slide 30: Plugins (continued)
NDPluginFFT
  Computes FFT of 1-D or 2-D NDArrays
  Exports NDArrays containing the absolute value (power spectrum) of the FFT
  Exports 1-D arrays of the FFT real, imaginary, absolute values, and time and frequency data.
NDPluginColorConvert
  Convert from one color model to another (Mono, RGB1 (pixel), RGB2 (row) or RGB3 (planar) interleave)
  Bayer conversion removed from this plugin, now part of Prosilica and Point Grey drivers.

## Slide 31: commonPlugins.adl All plugins at a glance

## Slide 32: ROI plugin

## Slide 33: Statistics plugin

## Slide 34: Statistics plugin (continued)

## Slide 35: Overlay plugin
Centroid of laser pointer calculated by statistics plugin
Cursor overlay X, Y position linked to centroid

## Slide 36: Processing plugin

## Slide 37: Processing plugin 30 microsec exposure time
No filtering
N=100 recursive average filter

## Slide 38: Transform plugin

## Slide 39: NDPluginPva (EPICS V4/7)
New plugin that converts NDArrays into the EPICS v4 normative type NTNDArray
Embedded EPICSv4 server serves the new NTNDArray structure as an EPICSv4 PV
High performance, ~3.2GB/s shown here
Can be received by any EPICS v4 client
  Java, Python, C++ versions of pvAccess
  CSS has a widget that can display NTNDArrays
  New ImageJ plugin
  Can include an NTNDArray receiver in another IOC
From Bruno Martins

## Slide 40: pvAccess Driver (EPICS V4)
Logical inverse of NDPluginPva
Receives NTNDArrays over the network, converts to NDArrays and calls plugins
Can be used to run areaDetector IOC and plugins on another machine or in another process
High performance:
  ~1.2 GB/s shown here with interprocess communication
  Saturating 10 Gb Ethernet links has been demonstrated
From Bruno Martins

## Slide 41: Plugins: NDPluginFile
Saves NDArrays to disk
3 modes:
  Single array per disk file
  Capture N arrays in memory, write to disk either multiple files or as a single large file (for file formats that support this.)
  Stream arrays to a single large disk file
For file formats that support it, stores not just NDArray data but also NDAttributes

## Slide 42: Plugins: NDPluginFile
File formats currently supported
  NDFileTIFF
    Supports any NDArray data type
    Stores NDAttributes as ASCII user tags
  NDFileJPEG
    With compression control
  NDFileNetCDF
    Popular self-describing binary format, supported by Unidata at UCAR
  NDFileHDF5
    Writes HDF5 files with the native HDF5 API, unlike the NeXus plugin which uses the NeXus API. Supports 3 types of compression.
    Supports using an XML file to define the layout and placement of NDArrays and NDAttributes in the HDF5 file
    Support Single Writer Multiple Reader (SWMR). Only supported on local file systems, GPFS, and Lustre (not NFS or SMB)

## Slide 43: Plugins: NDPluginFile
File formats currently supported
  NDFileNeXus
    Standard file format for neutron and x-ray communities, based on HDF5, which is another popular self-describing binary format; richer than netCDF
    May be deprecated in a future release since NeXus files can now be produced with the NDFileHDF5 plugin using an appropriate XML layout file
  NDFileMagick
    Uses GraphicsMagick to write files, and can write in dozens of file formats, including JPEG, TIFF, PNG, PDF, etc.
  NDFileNull
    Used only to delete original driver files when no other file plugin is running

## Slide 44: File saving with driver
In addition to file saving plugins, many vendor libraries also support saving files (e.g. marCCD, mar345, Pilatus, etc.) and this is supported at the driver level.
File saving plugin can be used instead of or in addition to vendor file saving
  Can add additional metadata vendor does not support
  Could write JPEGS for Web display every minute, etc.

## Slide 45: NDPluginFile display: TIFF
Example: saving 82 frames/second of 1024x1024 video to TIFF files, a few dropped frames.

## Slide 46: NDFileHDF5

## Slide 47: NDFileHDF5 XML file to define file layout
<xml>
<group name="entry">
<attribute name="NX_class" source="constant" value="NXentry" type="string"></attribute>
<group name="instrument">
<attribute name="NX_class" source="constant" value="NXinstrument" type="string"></attribute>
<group name="detector">
<attribute name="NX_class" source="constant" value="NXdetector" type="string"></attribute>
<dataset name="data" source="detector" det_default="true">
<attribute name="NX_class" source="constant" value="SDS" type="string"></attribute>
<attribute name="signal" source="constant" value="1" type="int"></attribute>
<attribute name="target" source="constant" value="/entry/instrument/detector/data"
type="string"></attribute>
</dataset>
<group name="NDAttributes">
<attribute name="NX_class" source="constant" value="NXcollection" type="string"></attribute>
<dataset name="ColorMode" source="ndattribute" ndattribute="ColorMode">
</dataset>
</group> <!-- end group NDAttribute -->
</group> <!-- end group detector -->
<group name="NDAttributes" ndattr_default="true">
<attribute name="NX_class" source="constant" value="NXcollection" type="string"></attribute>
</group> <!-- end group NDAttribute (default) -->
<group name="performance">
<dataset name="timestamp" source="ndattribute"></dataset>
</group> <!-- end group performance -->
</group> <!-- end group instrument -->
<group name="data">
<attribute name="NX_class" source="constant" value="NXdata" type="string"></attribute>
<hardlink name="data" target="/entry/instrument/detector/data"></hardlink>
<!-- The "target" attribute in /entry/instrument/detector/data is used to
tell Nexus utilities that this is a hardlink -->
</group> <!-- end group data -->
</group> <!-- end group entry -->
</xml>

## Slide 48: Multiple Threads per Plugin
Support for multiple threads running the processCallbacks() function in a single plugin.
Can improve the performance of the plugin by a large factor.
  Linear scaling with up to 5 threads (the largest value tested) observed for most of the plugins that now support multiple threads.
Maximum number of threads that can be used for the plugin is set in constructor and in IOC startup script.
Actual number of threads to use controlled via an EPICS PV at run time, up to the maximum value.
Optional sorting of NDArrays by uniqueId to attempt to output them in the correct order.
  Several new parameters to control this option
Plugins needed minor modifications to be thread-safe for multiple threads running in a single plugin object.
Most compute-intensive plugins now support multiple threads.

## Slide 49: Multiple Threads per Plugin 1 Thread

## Slide 50: Multiple Threads per Plugin 3 Threads

## Slide 51: Multiple Threads per Plugin 5 Threads

## Slide 52: NDPluginScatter
Used to distribute (scatter) the processing of NDArrays to multiple downstream plugins
  Allows multiple instances of a plugin to process NDArrays in parallel, utilizing multiple cores to increase throughput.
  Utilizes modified round-robin for choosing next output plugin
More complex than multiple threads in a single plugin, but allows the plugins running in parallel to have different configurations or even be different plugins
NDPluginGather
Merges NDArrays from multiple upstream plugins into a single output stream.
Designed to work with NDPluginScatter
Optional sorting by uniqueId

## Slide 53: Distributed Processing with NDPluginScatter + EPICS V4
Distribute HDF5 file writing to multiple IOCs (4096 x 3078 8-bit)
# IOCs | Files/sec | GB/sec
1 | 101.0 | 1.19
2 | 195.2 | 2.29
3 | 217.5 | 2.55
NDPluginPva
NDPluginPva
NDPluginPva
pvaDriver
NDFileHDF5
pvaDriver
NDFileHDF5
pvaDriver
NDFileHDF5
ADSimDetector
NDPluginScatter
IOC #1
IOC #2
IOC #3
IOC #4

## Slide 54: Viewers
areaDetector allows generic viewers to be written that receive images as EPICS waveform records over Channel Access
Current viewers include:
  ImageJ plugin EPICS_AD_Display. ImageJ is a very popular image analysis program, written in Java, derived from NIH Image.
  EPICS_NTNDA_Viewer. Same as above but uses pvAccess rather than Channel Access.
  ffmpegServer allows image display in any Web browser
  ffmpegViewer high-performance Qt-based viewer for MJPEG stream

## Slide 55: EPICS_NTNDA_Viewer ImageJ plugin
New ImageJ plugin written by Tim Madden and Marty Kraimer
Essentially identical to EPICS_AD_Viewer.java except that it displays NTNDArrays from the NDPluginPva plugin, i.e. using pvAccess to transport the images rather than NDPluginStdArrays which uses Channel Access.

## Slide 56: NDPluginPva Advantages
NTNDArray data transmitted "atomically" over the network
  Channel Access requires separate PVs for the image data and the metadata (image dimensions, color mode, etc.)
With Channel Access data type of waveform record is fixed at iocInit, cannot be changed at runtime.
  If the user wants to view both 8-bit images, 16-bit images, and 64-bit double FFT images then waveform record needs to be 64-bit double, adding a factor of 8 network overhead when viewing 8-bit images.
  pvAccess changes the data type of the NTNDArrays dynamically at run-time, removing this restriction.
Channel Access requires setting EPICS_CA_MAX_ARRAY_BYTES
  Source of considerable confusion and frustration for users.
  pvAccess does not use EPICS_CA_MAX_ARRAY_BYTES and there is no restriction on the size of the NTNDArrays.

## Slide 57: Viewers
EPICS_AD_Controller. Allows using the ImageJ ROI tools (rectangle and oval) to graphically define the following:
  The readout region of the detector/camera
  The position and size of an ROI (NDPluginROI)
  The position and size of an overlay (NDPluginOverlay)
  The plugin chain can include an NDPluginTransform plugin which changes the image orientation and an NDPluginROI plugin that changes the binning, size, and X/Y axes directions. The plugin corrects for these transformations when defining the target object.
  Chris Roehrig wrote an earlier version of this plugin.

## Slide 58: Other Drivers that use ADCore
NDArrays are not limited to 2-D detectors
  File, ROI, and statistics plugs are useful for other types of detectors
Used for spectra arrays [NumMCAChannels, NumDetectors, NumPixels] for:
  Xspress3 from Quantum Detectors
  xMAP, Mercury and new FalconX from XIA
Used for time-series data [NumTimePoints, NumInputs] for the quadEM quad electrometer software
  AH401, AH501, TetrAMM from CaenEls
  Two types of electrometers from BNL Instrumentation group (Peter Siddons)

## Slide 59: Conclusions
Architecture works well, easily extended to new detector drivers, new plugins and new clients
Base classes, asynPortDriver, asynNDArrayDriver, asynPluginDriver actually are generic, nothing “areaDetector” specific about them.
They can be used to implement any N-dimension detector, e.g. the XIA xMAP (16 detectors x 2048 channels x 512 points in a scan line)
Can get documentation and pre-built binaries (Windows) from our Web site:
  http://cars.uchicago.edu/software/epics/areaDetector
Can get code from github
  https://github.com/areaDetector