# D-DIA: A New Apparatus for High Temperature Triaxial Deformation to Pressures up to 15 GPa

To be submitted to Rev. Sci. Instrum.,

Tentative authorship: W. Durham, I. Getting, Y. Wang, D. J. Weidner

8/15/02

## 1. Introduction

Quantitative studies of rheological properties of materials under high pressure have been limited due to technical difficulties.  For example, in earth sciences, quantitative experimental studies of rheological properties of rocks and minerals as well as deformation-induced microstructures have been mostly limited to conditions below 4 GPa.  High-resolution rheological measurements using the off-the-shelf apparatus (e.g., the Paterson apparatus) can be made only up to about 0.5 GPa (Paterson, 1990), corresponding to a depth of ~15 km in the earth. With modifications to apparatus size, sample assembly, and institution of a molten salt confining medium, a Griggs solid-medium deformation apparatus can be used for quantitative rheological studies (e.g., Borch and Green), but the maximum pressure of its operation is still limited to about 4 GPa (Bai and Green, 1998; Renner et al., 2001), corresponding to a depth of about 120 km. As a consequence, rheological properties of more than 90% of Earth’s mantle remain unexplored in a quantitative manner.

Many novel techniques have been developed to address this need.  Kinsland and Bassett (1975) and Sun et al. (1977) used the diamond anvil cell (DAC) as a deformation apparatus.  In their studies, a sample is sandwiched between two single crystals of diamond and radial distribution of pressure is determined by measuring the shift of fluorescence lines of ruby crystals located at various points in the sample chamber. Assuming that the pressure gradient is supported by the sample strength, one can estimate the sample’s strength based on the equation of force balance. Such a technique has been used extensively by Meade and Jeanloz (1988, 1990) at room temperature to the pressure of ~40 GPa (see also Chai et al., 1998).  Major advantage of this technique is the high pressures that can be used under which a sample can be plastically deformed.  However, there are major limitations with this approach:  (1) homogeneous heating is difficult in a diamond anvil cell, and almost all previous results were obtained at room temperature, (2) the rate of deformation is not well constrained and deformation is likely to be time-dependent in these tests, and (3) the sample chamber is extremely small (typically microns to tens of microns) so that some important effects such as grain-size sensitivity of strength are difficult to measure.

There have been many efforts in using the multi-anvil press as a deformation machine.   In the multi-anvil press, compression of a sample occurs nearly isotropically.  However, by making the mechanical strength of a sample assembly anisotropic, one can create a deviatoric stress field in the sample assembly.  This deviatoric stress is relaxed by plastic flow upon heating.  This technique has been used in a number of studies (e.g., Fujimura et al., 1981; Fujimura, 1989; Green et al., 1990; Bussod et al., 1993; Karato and Rubie, 1997; Durham and Rubie, 1998; Karato et al., 1998). The analysis by Karato and Rubie (1997) (see also Durham and Rubie, 1998) showed that the mode of deformation in this technique is in most cases “stress relaxation”, i.e., the magnitude of deviatoric stress changes significantly within a single experiment.  The amount of piston movement in this technique is limited and hence the strain magnitude is limited.  However, by using a thin sample sandwiched between two pistons cut at 45, Karato and Rubie (1997) were able to attain relatively large strains (~1-2).  Quasi-constant displacement rate tests can also be made with the multi-anvil press through continuous movement of the ram (Green et al., 1990; Bussod et al., 1993).  Nevertheless, the maximum amount of strain attained by this mode of operation is limited because the confining pressure also changes with deformation and gasket extrusion eventually stops.  The multi-anvil deformation technique has been applied to ~25 GPa and ~2000 K with a sample dimension of ~1-2 mm diameter and ~2-4 mm long, or a thin (~0.2 mm) disk sample with a similar diameter (Cordier et al., 2001).  Major advantage of this technique as compared to the DAC technique is that rheological properties can be investigated under more homogeneous high temperature and pressure conditions with better-controlled chemical environment.  Also, because relatively large samples can be used, microstructural evolution during deformation and its effects on rheology can be investigated (e.g., Karato et al., 1998).  This technique has two major limitations, however.  First, pressurization and plastic deformation cannot be completely separated and deformation during pressurization likely occurs unless special care is taken in the sample assembly to absorb initial stage shape change of sample assembly (Karato and Rubie, 1997; Durham and Rubie, 1998).  Second, because the stress magnitude changes during an experiment, dominant mechanisms of deformation may change in a single run, making interpretation of mechanical data and microstructure difficult.

Yamazaki and Karato (2001) modified the Drickamer-type high-pressure apparatus to achieve high strain in the sample.  One of the anvils is fixed with the frame, whereas the other is attached to a rotational actuator.  The sample is twisted between the two pistons, and the geometry of deformation is simple shear.  Yamazaki and Karato reported that with their motor and gear combination, shear strain rates from ~0 to ~10-3 s-1 can be achieved.  Extremely large strains (in principle unlimited) can be attained.  However, it is difficult to quantify the stress-strain relationship.

To achieve controlled deformation under high pressures, Shimada (1981) deployed a cubic press, which drives six independent hydraulic rams to push six square-tipped anvils, compressing a cubic pressure medium.  To perform triaxial deformation tests, the upper and lower anvils are advanced independently while keeping the four horizontal anvils fixed.  However, precise manipulation of six independent hydraulic rams is both expensive and operationally complicated.  No new activities were reported after initial deployment of this apparatus to our knowledge.

In this paper, we report a new high pressure deformation apparatus that is capable of generating pressures up to 15 GPa (or about 450 km depth in the earth) at high temperatures (up to ~2000C).   This apparatus is a modification of the cubic anvil apparatus known as the DIA (e.g., Osugi et al., 1964; Shimomura et al., 1985), which is widely used in solid-media hydrostatic compression experiments coupled with synchrotron X-ray diffraction (Inoue and Asada, 1973; Shimomura et al., 1985; Weidner et al., 1991; Wang et al., 2000).  This apparatus takes advantages in recent developments using synchrotron radiation to allow us to measure displacement and stress to the levels of <1 m and ~10 MPa, respectively, thus making it possible to perform quantitative deformation experiments under pressure and temperature conditions corresponding to the transition zone.

2. Design Criteria

Configuration

The original DIA apparatus consists of two (upper and lower) pyramidal guide blocks or bolsters, four trapezoid end blocks (thrust blocks), and six anvil holders, as skematically illustrated in Figure 1. The inner surfaces of the guide blocks form a tetragonal pyramid. Two of the six anvils (upper and lower) are along the center line of this pyramid and are fixed opposite to each other on each guide block. The other four anvils are horizontally located on the midpoints of the square edges of a bipyramid. This results in the formation of a cubic nest bounded by the flat faces of the six anvils. The ram force is applied along the vertical axis. The six equivalent components of forces act along the three orthogonal axes to force advancement of the six anvils toward the center of the cube. The advance of all the anvils is automatically synchronized by careful alignment so that the cubic sample furnace assembly is compressed hydrostatically.

In our new apparatus called deformation DIA (or D-DIA), two additional hydraulic rams are introduced in the upper and lower bolster blocks.  These differential rams allow the top and bottom anvils to be advanced and retracted independently.  Figure 2 shows a cutaway view of the D-DIA concept.  The differential rams are located within the guide blocks of the D-DIA module and pushed against the platens driven by the main hydraulic ram, which is used to apply hydrostatic pressure to the sample chamber through solid pressure media.

Figure 1.  Conceptual diagram illustrating the principles of the DIA cubic anvil apparatus.  See text for detailed description.

### Load consideration

Typical main forces are up to ~100 ton in the hydrostatic DIA.  Maximum attainable pressures vary depending on the anvil truncation size.  When tungsten carbide (WC) anvils are used, the maximum pressures are about 7 GPa for 6 mm truncation anvils (sample cell length 8 mm), 10 GPa for 4 mm (6 mm cell), and 13 GPa for 3 mm (5 mm cell) (Shimomura et al., 1992; Utsumi et al., 1992), respectively.  Wang et al. (1996) showed that with a slight taper (4) near the anvil tip, the modified 4 mm anvils can reach pressures about 35% higher than the non-tapered anvils.  With sintered diamond anvils, pressures up to 18 GPa are achievable with 3 mm truncation (e.g., Utsumi et al., 1992).

Ignoring friction between all the sliding surfaces, each anvil will be pushed by a force that is 1/3 of the main ram force.  The differential rams are capable of generating 70 Ton of force (at 100 MPa oil pressure in the differential rams), sufficient to overcome the confining force and apply differential force even at the highest confining pressure.

Differential Rams

Figure 2.  Cutaway view of the D-DIA with differential rams in the upper and lower guide blocks. Note locations of the differential rams.

A typical sample length is 1 to 2 mm.  So a total differential displacement of 1 mm would result in 50% deformation, if all the displacement is transmitted to the sample.  However, as all the cell parts will deform, only a portion of the strain will be partitioned into the sample.  Therefore, our goal is to have a maximum piston displacement of 2 mm in each differential ram, to ensure sufficient deformation in the cell.

The differential rams are designed by considering the space limitation inside the guide block, the maximum oil pressure in the ram, and the maximum piston displacement requirement.  The resulting ram has an inner diameter of 89.00 mm, with a piston that has a tapered cross section (Fig. 3).  The cross section of the bore that is cut out from the guide block has the same tapered angle, hence limiting the piston travel to 2 mm.  A backing plate is used to close out the oil chamber.  Both the piston and the backing plate have a ledge for the O-ring assembly.  Sealing is achieved by Parker (N0552-90) O-rings (Nitrile, 90 durometer A), with Parker Parbak N0300-90 back-up rings.

Figure 3.  Details of the differential ram design.

### X-Ray Access

A very important requirement for the D-DIA is the ability to perform diffraction and imaging analyses on the sample using synchrotron X-radiation.  The original hydrostatic DIA has been widely used for synchrotron applications (e.g., Shimomura et al., 1985; Weidner et al., 1992; Wang et al., 2000).   However, in the past, diffraction has been limited to the vertical plane only.  This is largely due to the anvil material of choice, namely WC, which is highly attenuating to X-rays.  In order to allow precise stress measurements using the X-rays, distortion in the Debye rings must be measured accurately.  This requires multiple diffractions at various angles in the plane perpendicular to the incident X-rays.  To facilitate this, conical openings are made into the guide blocks, so that a 10 solid angle can be accessed (Figure 3).  Notches are made in the anvil-supporting wedges to give access to the entire 10 solid angle.  With the advent of new anvil materials such as sintered diamond or cubic boron nitride (cBN), this configuration will allow us to measure complete Debye rings for stress determination.

## 3. Finite Element Analyses

Three-dimensional finite element analyses (FEA) were performed on the guide blocks to examine impacts on the mechanical integrity under the maximum main ram and differential ram load, after introduction of the differential ram.  The results indicated that under a main ram load of 1 MN (vertical) and a differential load of 105 MPa (inside the bore for the differential pistons), the maximum Von Mises stress is ~ 370 MPa.   The guide block deforms very little: the maximum horizontal displacement, which occurs on the base of 45 sliding surfaces, is about 10 µm.   Overall the largest displacement magnitude is 60 µm.

Two dimensional, axial symmetrical FEA was carried out on the differential ram itself.  Figure 4a shows Von Mises stress magnitude in the differential ram piston.  The piston is a combination of a cylinder and a cone.  Rotating the cross section in Fig. 4a by 360 will complete the shape.  The maximum stress (877 MPa) that occurs at the sharp corners in the piston is clearly an artifact.  In reality, the corners are rounded with a 2 mm radius and the stress concentration should disappear.  Away from the corners, the highest Von Mises stress is in 300 MPa level.

Figure 4a also shows exaggerated displacement field.  The maximum displacement occurs in the lower right corner and is 40 µm and almost purely vertical.

Figure 4a.  Von Mises stress and deformation of the differential ram piston.

Similar FEA results on the ram base plate and the guide block are summarized in Figures 4b and c.  For the ram base plate, maximum Von Mises stress is in the order of 120 MPa if corner effect is removed and maximum displacement is only 7 µm in the upper right corner.

Figure 4b. Von Mises stress magnitude in the ram base plate.  Maximum stress occurs at the lower right corner, in the order of 120 MPa.  Maximum displacement is 7 m, occurs near the upper right corner.

Figure 4c.  Displacement and Von Mises stress in the guide block, which is approximated as axial symmetric.  Maximum Von Mises occurs on the inner wall of the ram, around140 MPa.  Maximum displacement about 20 m occurs at the same location, almost pure radial.

## 4. Mechanical Setup

### Guide Blocks

Figure 5a shows the actual DDIA guide blocks, which are mounted in a die set to allow easy transportation in and out of the hydraulic press used for pressure generation.  The die set has four ball-bearing rollers to allow the entire unit to be rolled in and out of the press on rails.  A vertical LVDT is used to monitor closure between the top and bottom guide blocks during the experiment, and a horizontal LVT is mounted on the back of one of the four anvil bearing wedges to record travels of the horizontal anvils during deformation.  Four miniaturized displacement transducers (DVRTs from xxx company) are used to measure differential ram displacements (Figure 5b).

A

B

Figure 5.  Actual DDIA tested.  A is Overall view of the setup and B shows details of the DVRTs used for monitoring differential ram displacement during deformation

### Hydraulic Control

As the differential rams are located within the D-DIA module and will push against the platens driven by the main hydraulic ram, independent hydraulic control is required for deformation.  Hydrostatic pressure is generated to the same levels and in the same manner as with the hydrostatic DIA.  To apply a deformation field under pressure, the upper and lower anvils need to be pushed simultaneously towards the center of the sample assembly.  To maintain the tetragonal symmetry, the two differential rams need independently controlled with ram displacements closely monitored and adjusted so that they travel by the same amount.  In addition, any change in pressure in the differential rams will change the main ram pressure accordingly; therefore some fine control on the main ram is needed.

We use three positive displacement pumps (PDP) to achieve this goal.   Two PDPs are used to control differential ram displacement, and the third is tapped into the main hydraulic ram on the hydraulic press to allow finer control in oil pressure.   Each pump has an oil capacity of ~10 cc  and its displacement can be accurately controlled.  Brand and specs of the PDPs

## 5. Test Results using Synchrotron X-ray Sources

To perform quantitative deformation experiments in the D-DIA, one needs to be able to measure both stress and strain under high pressure.  Conventional load cell and strain gage techniques are no longer applicable in such solid medium apparatus.  With the advent of synchrotron X-ray techniques, two well-established synchrotron X-ray detection techniques can be used for these purposes.  The D-DIA can readily be fitted to synchrotron radiation facilities to allow quantitative measurements of stress and strain at high-pressure and temperature conditions.

Figure 6 shows a typical 6-mm cell assembly used for the tests.  Both pyrophyllite and boron-epoxy mixture were used as pressure media, with a thermocouple entering the cylindrical graphite heater (4 mm long) from the top.  Temperature gradients, estimated from unit cell volume measurement on NaCl along the axial direction, were about 200C within the center 2.0 – 2.5 mm and within 50C over a 1 mm length near the center.  Crushable alumina (porosity about 30%) end plugs were used to minimize axial deformation during cold compression.

Figure 6. Cell assembly used in deformation experiments.  Total sample length is typically 2 mm, located in the center of the graphite heater.

### Sample Length and Strain Measurement

Synchrotron X-rays can also be used to measure sample length. Direct images of markers can be obtained using an incident beam whose dimensions are larger than the sample and a marker with contrasting absorption characteristics. Platinum or gold foils above and below the sample are easily viewed. The X-ray image is obtained by projecting the X-ray onto a fluorescent screen that is viewed with a CCD camera though a magnifying system (Wang et al., 2000；Chen et al., 2001).  By analyzing the images during the deformation, sample length change and strain can be determined.   Another technique is to detect fluorescence of the marker material used in the cell to define the ends of the sample (Canistraro et al., 1998).  In either case, it has been demonstrated that a sample length can be measured to within 0.1 m.  With a sample length of 2 mm, this means that strains of 10-4 or better can be measured. For images taken 100 seconds apart, this allows resolution of strain rates of 10-6/s.

The DDIA has been successfully tested both at X-17B1 using SAM-85 (Weidner et al., 1992) at the National Synchrotron Light Source (NSLS) and the 250 ton press in 13-BM-D at GSECARS (Rivers et al., 1996).  Figure 7 is a screen dump showing the vital information being monitored and controlled for the DDIA during deformation of a fully densified polycrystalline Mo sample.  The differential ram displacements were monitored by the four DVRTs as well as ram hydraulic pressures.  During deformation both differential rams were advanced at a controlled rate of about 1.5x10-5 mm/s, resulting in a total cell shortening rate of 3x10-5 mm/s.  The differential ram pressures (PT1 and PT2) initially needed to increase to achieve this rate.  The advance of the differential rams also caused the main ram pressure (PT3) to increase.  In order to maintain a constant cell pressure, the main ram pressure was decreased.  From this point on, the differential rams were at a more or less constant pressure.

Both the vertical guide block gap and the horizontal anvil displacement were also monitored.  The observations show that as the differential rams were advanced and main ram pressure reduced, vertical guide block gap continued to decrease, confirming that the cell was indeed shortened continuously.  Also as the top and bottom anvils advanced, the horizontal anvils retracted, as confirmed by the LVDT monitoring the horizontal wedges.  All these observations confirmed that the DDIA was behaving as expected and capable of producing up to 1 mm shortening in the cell assembly.

Figure 7. Example of control record showing machine parameters during a deformation experiment (D0374).  DVRT1 and 2 are the transducers monitoring displacement on the top differential ram; DVRT5 and 7 for the bottom differential ram.  Amplitudes are in volts (1 V ~ 0.5 mm).   Pressures in the differential rams are monitored by the pressure transducers, PT1 and 2, with calibration constant 1 V~ 6000 psi. LVDT1 and 2 are for gap between guide blocks and horizontal anvil outward displacement, respectively (calibration constants: 1 V ~ 2.54 mm).

Figure 8 shows sample length as a function of time during this experiment, where a fully densified polycrystalline Mo sample (1 mm long) was deformed back-to-back with a mixture of Au and MgO at 5 GPa and 1000C.  Sample length change was measured from digital X-ray imaging.  Within about 5 hr, the sample was shortened axially by 0.2 mm (about 20%).  Clearly, cold compression did not introduce large amount of deformation, due to the cushioning introduced by the crushable alumina end plugs and the MgO+Au pressure standard.  However we pay a price for this cushsioning: out of the 0.4 mm total shortening (as indicated by the DVRT records), only 50% was transmitted to the sample.  Clearly, some work is needed in optimize the sample design in order to maximize deformation while maintaining relatively hydrostatic environment during cold compression.

X-ray imaging also confirmed the horizontal anvil displacement observations from the LVDT measurement.  As the differential rams continued to push the cell in the vertical direction, horizontal anvils moved outward, causing the vertical anvil gaps to increase.  Figure 9 also shows the measured anvil gap change during deformation.  The rate of gap opening is in the same order of magnitude as the sample shortening in this experiment.

Figure 8.  Deformation of a polycrystalline Mo sample (1 mm) at 5 GPa and 1000C.  Upper panel shows sample length measurement, which gives a strain rate of 10-5 s-1.  Lower panel shows associated horizontal gap opening.

In a separate experiment, two samples, polycrystalline rods of MgO and Ta were “piggybacked” in one cell assembly and deformed under a hydrostatic pressure of 10 GPa at 500C.  The pressure was measured based on the equation of state of MgO (Utsumi et al., 1992) and sample lengths of MgO and Ta were determined using X-ray imaging.  Figure 9 shows strains measured from the two samples under those conditions.  Evidently Ta and MgO have similar flow properties.  This is the portion of the data collected at 10 GPa and 500C.  Pressure and temperature are both held constant during the deformation.

Figure 9.  Total axial length change of the two polycrystalline samples (MgO and Ta) deformed at 10 GPa and 500C.  Note the samples deform at different rates, especially during the first deformation step (1.2 vs. 0.8x10-5 s-1 for MgO and Ta, respectively), after which, the MgO sample shortened more substantially while the differential piston advancement was stopped (the waiting period between the two deformation steps).  During the second deformation step, the strain rate of MgO is only slightly faster than that of Ta.

### Stress Measurement

The technique for measuring elastic strain using X-ray diffraction to obtain information on stress in materials dates back at least to the 1930’s (e.g., see review by Greenough, 1952; also Bragg, 1942; Stoke et al., 1943).  Applications to high pressures in situ have become feasible using high-energy X-rays generated by a synchrotron source (for a review see Weidner, 1999).  X-rays diffractions measure the distance between lattice planes (called d-spacing) of the crystalline sample. Plastic strains are macroscopic shape change of a sample, achieved by  This distance is insensitive to strain history (plastic strain), but reflects the elastic strain field.  In non-hydrostatic stress field, crystallographic planes with different orientations change their spacing (d-spacing) to a different degree.  For polycrystalline samples where grains are randomly oriented, each spacing is seen as a continuous ring (so-called Debye ring) from X-ray diffraction.  In the presence of non-hydrostatic stress, the d-spacing of a given crystallographic plane will depend on the orientation of crystal with respect to the applied stress and the stress magnitude.  Thus the Debye ring will be distorted.  By measuring the orientation dependence of d-spacing (or the distortion of the Debye ring), one can determine the stress magnitude using the known elastic constants of the crystal.  The anvils themselves cast a shadow and limit the region where the diffracted signal can be observed.  However, with the advent of anvil materials, especially the cBN anvils, which are X-ray transparent, it is now possible to observe the entire Debye ring.  This measurement can now be done in a multianvil high-pressure apparatus using a synchrotron X-ray source (Weidner et al., 1992; Vaughan et al., 1998).  With a precision of 10-4 for the lattice spacing and a typical Young’s modulus of 200 GPa it should be possible to resolve differential stresses of ~20 MPa.  Analysis of the effect of a uniaxial stress field for anisotropic crystals has been discussed by Singh (1993) and by Singh and Balasingh (1994).

Although stress measurements are based on elastic distortion in a crystalline lattice, it does not necessarily mean that we have to invoke an elasto-plastic model in measuring stress-strain curves for a given sample.  As the sample of interest can be stacked with other “reference” materials in the deformation column, one can carefully select a “piezometer”, which, in ideal situation, ought be inside the elastic region when the sample deforms non-elastically.  The stress information obtained from the “piezometer” can then be used to infer the stress level in the sample.

### Other Modes of Experimentation

Once the cubic symmetry of the hydrostatic DIA is relaxed to tetragonal in the D-DIA and the coupling between hydrostatic pressure and deformation is removed, there are many new ways to perform deformation experiments.  One obvious variation is to perform tensile experiments under pressure as well as compression.  This can be achieved with the differential rams extended before loading the sample hydrostatically.  Once the sample is under pressure, the differential rams can be retracted and the advanced again.  Such tensile – compression cycles can be repeated in a single run.  This is possible because regarding complex deformation of gaskets, Wang et al. (1998) successfully demonstrated that pressure can be maintained hydrostatic in the DIA despite repeated reverse motion on the pyrophyllite gaskets.

Another possible mode is to use anvils of different shapes in order to optimize gasket performance. With ordinary 45°-cut tungsten carbide anvil, the eight end-anvil gaskets will sense compression (pure shear) during deformation. Sample loading efficiency may be poor in this circumstance because the gaskets must be continuously thinned. With end (upper and lower) anvils having an cut 26.6° (arctan 0.5), volume-compensating motion of end and side anvils will be such that end-anvil gaskets will have constant thickness as end anvils advance. Loading efficiency will be improved, and larger deformations will be possible.

## 6. Conclusions

The DDIA apparatus is based on the DIA with the following modifications: (1) Introduction of two additional hydraulic (differential) rams inside the upper and lower bolster blocks in order to drive the upper and lower anvils independently; (2) the two differential rams are independently controlled, so that both the upper and lower anvils can be advanced or retracted independently in order to maintain the mirror symmetry between the upper and lower guide blocks; (3) an additional (fine) hydraulic control is added to the main ram that is used to compress the entire D-DIA assembly, in order to maintain desired pressure as the differential anvils are advanced or retracted.  The independent control of confining and differential rams means that deformation can be imposed without the necessity of increasing confining pressure (which has been the longstanding limitation of deformation studies in MA 6/8 multi-anvil and in diamond-anvil assemblies).

We have demonstrated that the D-DIA works.  Up to 20% axial deformation has been generated on a 1 mm long sample at high pressure and temperature conditions of 10 GPa and 1200C, respectively.  Temperature and hydrostatic pressure are generated to the same levels and in the same manner as with the conventional DIA.  The differential rams are capable of generating 70 T force, sufficient to overcome the confining force and apply differential force even at the highest confining pressure.  Differential ram velocities can be controlled from approximately 10-7 to 10-2 mm/s.  Assuming balanced motion of the differential rams is the appropriate mode of operation, this translates to a strain rate on typical a 1-mm-length, 1-mm3-volume sample of 2 x 10-2  /t  2 x 10-7 s-1.

Acknowledgments: Design: Paul Murry, Clayton Pullin, and the CARS drafting team; Setup: D. Ruddle, Carey Koleda, Bayreuth machine shop; Testing: S. Mei, T. Uchida, Dan Frost, Jiuhua Chen, Mike Vaughan, Li Li, Pamela Burnley.

## References

Borch, R. S. and H. W. Green, II. 1987 Dependence of creep in olivine on homologous temperature and its implications for flow in the mantle. Nature 330, 345-348.

Bragg, L. (1942) A theory of the strength of metals, Nature, 149, 511-513.

Bussod, G.Y., Katsura, T. and Rubie, D.C., 1993. The large volume multi-anvil press as a high P-T deformation apparatus, PAGEOPH, 141, 579-599.

Canistraro, HA, EH Jordan, DM Pease, (1998) Differential displacement measurement using scanning x-ray beams, Rev. Sci. Instrum., 69, 452-456.

Chai, M., Brown, J.M, and Wang, Y., 1998. Yield strength, slip systems and deformation induced phase transformation of San Carlos olivine up to the transition zone pressure at room temperature, in Earth and Planetary Materials at High Pressure and Temperature, Geophysical Monograph 56, ed. by M. Manghnani and T. Yagi, AGU, Washington, D. C., pp. 483-493.

Cordier, P. and Rubie, D.C., 2001. Plastic deformation of minerals under extreme pressure using a multi-anvil apparatus, Mater. Sci. Eng., A309-310, 38-43.

Davies, G.F., 1995. Penetration of plates and plumes through the mantle transition zone, Earth Planet. Sci. Lett., 133, 507-516.

Durham, W. B. and D. C. Rubie, 1998. Can the multianvil apparatus really be used for high pressure deformation experiments?, in Earth and Planetary Materials at High Pressureand Temperature, Geophysical Monograph 56, ed. by M. Manghnani and T. Yagi, AGU, Washington, D. C., pp. 63-70.

Fukao, Y., Obayashi, M., Inoue, H. and Nenbai, M., 1992. Subducting slabs stagnant in the transition zone, J. Geophys. Res., 97, 4809-4822.

Fujimura, A., 1989. Preferred orientation of mantle minerals, in Rheology of Solids and of theEarth, (eds. S. Karato and M. Toriumi), Oxford University Press, pp. 263-283.

Fujimura, A., Endo, S., Kato, M. and Kumazawa, M., 1981. Preferred orientation of _-Mn2GeO4, Programme and Abstracts, Seism. Soc. Japan, pp.185.

Funamori, N. and Yagi, T. (1993) High pressure and high temperature in situ X-ray observation of MgSiO3 perovskite under lower mantle conditions, Geophys. Res. Lett., 20, 387-390.

Greenough, G.B. (1952) Quantitative X-ray Diffraction Observations on Strained Metal Aggregates, in: Progress in Metal Physics, Vol. 3, ed. Chalmers, B., Interscience Publishers, New York, pp. 176-219.

Green (1994);

Griggs and Baker, 1969

Inoue, K. and T. Asada (1973) Cubic anvil X-ray diffraction press up to 100 kbar and 1000C, Jpn. J. Appl. Phys., 12, 1786-1793.

Karato, S., 1998. Effects of pressure on plastic deformation of polycrystalline solids: some geological applications, Mat. Res. Soc. Symp. Proc. vol 499: 3-14.

Karato, S. and Rubie, D.C., 1997. Toward an experimental study of plastic deformation under deep mantle conditions: A new sample assembly for deformation experiments under high pressures and temperatures, J. Geophys. Res., 102, 20,111-20,122.

Karato, S., Dupas-Bruzek, C. and Rubie, D.C., 1998. Plastic deformation of (Mg,Fe)2SiO4 spinel (ringwoodite) under the transition zone conditions, Nature, 395, 266-269.

Kinsland, GL and WA Bassett (1975) Strength of MgO and NaCl polycrystals to confining pressures of 250 kbar at 25 C, J. Appl. Phys., 48, 978-985.

Meade, C. and Jeanloz, R. (1988) Yield strength of B1 and B2 phases of NaCl, J. Geophys. Res., 93, 3270-3274.

Meade, C. and R. Jeanloz (1988)

Meade, C. and Jeanloz, R. (1990) The strength of mantle silicates at high pressures and room temperature: implications for the viscosity of the mantle, Nature, 348, 533-535.

Moller, H., J. Barbers (1935), J. Mitt. Kaiser Wilhelm-Inst. Eisenforsch, 17, 81-

Morris, S., 1995. Stress relief during solid-state transformation in minerals. Proc. R. Soc. London, Ser. A 436, 203-216.

Osugi, J., K. Shimizu, K. Inoue, and K. Yasunami (1964) A compact cubic anvil high pressure apparatus, Rev. Phys. Chem. Jpn., 34, 1-6.

Paterson, M.S., 1990. Rock deformation experimentation, in The Brittle-Ductile Transition inRocks, (eds. A.G. Duba et al.), AGU, pp. 187-194.

Paterson, M.S. and Olgaard, D.L., 2000. Rock deformation tests to large shear strains in torsion, J. Struct. Geol., 22, 1341-1358.

Renner et al., 2001. An experimental study into the rheology of synthetic polycrystalline coesite aggregates. J. Geophys. Res. (in press).

Rubie, D.C., Karato, S., Yan, H. and O’Neill, H.C.St., 1993. Low differential stress and controlled chemical environment in multianvil high-pressure experiments, Phys. Chem.Mineral., 20, 315-322.

Shimada, M (1981) The method of compression test under high pressures in a cubic press and the strength of granite, Tectonophysics, 72, 343-357.

Shimomura, O., S. Yamaoka, T. Yagi, M. Wakatsuki, K. Tsuji, H. Kawamura, N. Hamaya, O. Fukuoga, K. Aoki, and S. Akimoto, Multi-anvil type X-ray system for synchrotron radiation, in: Solid State Physics Under Pressure, ed. S. Minomura, Terra Scientific Publishing Co., Tokyo, 1985, pp. 351-356.

Shimomura, O., W. Utsumi, T. Taniguci, T. Kikegawa, T. Nagashima (1992) in: High-pressure research: Application to Earth and planetary sciences, ed. Y. Syono, M. H. Manghnani, Terra Scientific / American Geophysical Union, Washington, DC,  pp. 3-11.

Singh, AK (1993)  The lattice strain in a specimen (cubic system) compressed nonhydrostatically in an opposed anvil device, J. Appl. Phys., 73, 4278-4286.

Stokes, A. R., K. J. Pascoe, H. Lipson (1943) X-ray evidence of the nature of cold work in metals, Nature, 151, 137.

Sung, C.M., C. Goetze, and H.K. Mao (1977) Pressure distribution in the diamond anvil press and the shear strength of fayalite, Rev. Sci. Instrum., 48, 1386-1391.

Tingle, TN, HW Green, TE Young and TA Koczynski (1993) Improvements to Griggs-type apparatus for mechanical testing at high pressures and temperatures, PAGEOPH, 141, 523-543.

Tullis, T.E. and Tullis, J., 1986. Experimental rock deformation techniques, in Mineral and Rock Deformation: Laboratory Studies—The Paterson Volume, (eds. B.E. Hobbs and H.C. Heard), AGU, pp. 297-324.

Uchida, T, N. Funamori, and T. Yagi (1996) Lattice strains in crystals under uniaxial stress field, J. Appl. Phys., 80, 739-746.

Utsumi, W., T. Yagi, K. Leinenweber, O. Shimomura, T. Taniguchi (1992) High pressure and high temperature generation using sintered diamond anvils, in: High-pressure research: Application to Earth and planetary sciences, ed. Y. Syono, M. H. Manghnani, Terra Scientific/American Geophysical Union, Washington, DC,  pp. 37-42.

Vaughan, M. T., D. J. Weidner, Y. B. Wang, J. H. Chen, C. C. Koleda, and I. C. Getting (1998) T-CUP: A new high-pressure apparatus for X-ray studies, The Review of High Pressure Science and Technology 7, 1520-1522

Wang, Y., Getting, I.C., Weidner, D.J., and Vaughan, M.T. (1998) Performance of tapered anvils in a DIA-type, cubic-anvil, high-pressure apparatus for X-ray diffraction studies, in Properties of Earth and Planetary Materials, Geophysical Monograph 101, ed. by M. Manghnani and T. Yagi, AGU, Washington, D. C., pp. 35-39.

Wang, Y., M. Rivers, T. Uchida, P. Murray, G. Shen, S. Sutton, J. Chen, Y. Xu, D. Weidner (2000), High pressure research using large-volume presses at GeoSoilEnviroCARS, Advanced Photon Source, in: Proceedings of the International Conference on High Pressure Science and Technology (AIRAPT-17), Science and Technology of High Pressure, vol. 2, pp. 1047-1052.

Weidner, D.J., Rheological studies at high pressure, in Reviews in Mineralogy, edited by P.H. Ribbe, pp. 493-524, Mineralogical Society of America, Washington D.C., 1998.

Weidner, D. J., M. T. Vaughan, J. Ko, Y. Wang, X. Liu, A. Yeganeh-Haeri, R. E. Pacalo, Y. Zhao (1992)  Characterization of stress, pressure, and temperature in SAM85, A DIA type high pressure apparatus. In: High-Pressure Research: Application to Earth and Planetary Sciences. Eds. Y. Syono and M. H. Manghnani. Tokyo and Washington, D.C., Terra Scientific Publishing Company and American Geophysical Union, pp. 13-17.

Yamazaki, D., and S. Karato (2001) High pressure rotational deformation apparatus to 15 GPa, Rev. of Sci. Instrum., 72, 4207-4211.